## Revision Control
Created by: |  |  | Program: VT BSME Senior Capstone
Created Date: |  |  | Project: 205 Assistive Robotic Arm for PPMD
Rev # | Edited By | Edited Date | Description of addition and/or change
1 | AJ | 2024-09-11 | Populating document with general requirements regarding robotic aid and expected functions
2 | AJ, Brooke, Cavan, Justin, Omar | 2024-09-17 | Refinement of requirements focusing on shalls and shoulds
3 | Brooke | 2024-09-18 | Refinement of requirements after feedback from sponsor

## Requirement_Traceability_Matrix
REQUIREMENTS VERIFICATION TRACEABILITY MATRIX (RVTM) | [7 empty] | Legend
Project Name: | Robotic Arm to Assist Motorized Chair Users |  | Customer Approval |  |  |  |  | TR | 1.0 | Overall configuration, performance specification, and description
Program/Customer: | PPMD | [6 empty] | TR | 2.0 | Field use conditions and environmental effects mitigation
Project Manager Name: | Brooke Harrington |  | Signature Keith Van Houten | 09/19/2024 |  |  |  | TR | 3.0 | Maintenance, installation, service life
Project Description: | Robotic arm that attaches to a motorized chair to aid people with mobility restrictions due to Duchenne Muscular Dystrophy | [6 empty] | TR | 4.0 | Physical features and metrics
Technical Requirement ID # | Requirement | Requirement Origination | Comments | Verification Method | Test Case Number | Verification | Status | TR | 5 | Human and robot safety
TR 1.0.0 | The team shall research, design, develop, integrate, and test an assistive robotic device modified for a motorized wheelchair including a custom mounting device, electronics, and software. | Customer Requirement | The robot should be able to preform one well defined task. More task can be undertaken as the main task is completed.
TR 1.1.0 | The robot shall be controlled manually by a joystick and should have the ability to program automatic functions accessed through a button push | Customer Requirement | The control system should be able to attach to either left/right side (dependent on user's dominant side) | Demonstration |  |  | In Progress | O | 1
TR 1.1.1 | The joystick shall be attached in a location near the motorized wheelchair armrest | Customer Requirement | Customer stated postioning the control interface near the non dominant hand so the wheelchair and robot system can be used simulataneously | Visual Inspection |  |  | Met with current design
TR 1.1.2 | Any automatic functions shall be independent of the joystick and should operate without any joystick inputs | Customer Requirement | Automatic functions are to be determined if main function is properly adressed | Program Inspection
TR 1.1.3 | The arm shall not mount in a location that infringes on the user's seating space, and must be small enough that as installed on wheelchair, the user can still pass through a standard door opening | Customer Requirement |  | Physical Inspection |  |  | Met with current design
TR 1.1.4 | The arm should be able to bring a drink of weight 470g to a specified location in front of the user | VT Design Team |  | Test |  |  |  | B | 5
TR 1.2.0 | The robot shall have no more than six degrees of freedom | Customer Needs | Developed from "ease of use" customer need | Inpection |  |  | Met with current design
TR 1.3.0 | The robot shall have an operating radius that spans a minimum of 5 inches to a maximum of 2 feet from mount location with 180 degrees of rotation total | Customer Needs | Developed from high value functions discussion with customer | System Test |  |  | Met with current design
TR 1.3.1 | The device should be able to reach and depress elevator buttons and the handicap door button between 34" and 48" as specified by ADA standards | Customer Needs | Developed from researching ADA standards with button placement | Demonstration |  |  | In Progress | O | 2
TR 1.3.2 | The system shall have an accuracy of +-10mm | Product Quality | Will be both by analysis and by test | Analysis, Test |  |  | Analysis in progress, test in spring | O | 3
TR 1.4.0 | The robot and necessary electronics shall receive power from a rechargeable battery independent of the wheelchair electronics | Customer Needs |  | Inspection |  |  | Met with current design
TR 1.4.1 | The battery that powers the robot arm shall last 24 hours between charges | Customer Needs |  | Analysis |  |  | Met with current design
TR 1.4.2 | The battery shall remain installed in the case during recharge and should only be removed for maintenance or installation | Customer needs |  | Inspection |  |  | Met with current design
TR 2.0.0 | The robot shall operate in dry, arid conditions and should operate under rainy, humid conditions | Customer Needs |  | Demonstration |  |  | In Progress
TR 2.1.0 | The robot electronics shall be contained in a weather resistant enclosure | Customer Needs | Devloped from 2.0.0 | Inspection |  |  | In progress
TR 2.1.1 | The robot motors shall be secured in a weather resistant sleeve | Customer Needs | Devloped from 2.0.0 | Demonstration |  |  | In progress
TR 3.0.0 | The robotic aid shall be installed in less than one hour and should be operational within thirty minutes upon installation | VT Design Team | In order for multiple people from PPMD to test the system when they visit, it is ideal to have the install/unsinstall and time till operation be minimized without sacrificing safety or quality (subject to change once robot and assistive wheelchair have been acquired) | Timed Test |  |  |  | J | 6
TR 3.1.0 | The robot shall be disassembled and reassembled for maintenance in less than three hours OR The system shall have maintenance work, including assembly and disassembly, completed in less than three hours | VT Design Team |  | Timed Test or PM/M Plan Development |  |  |  | A | 7
TR 3.2.0 | The robot shall require maintenance every year and should be cleaned and inspected every 3 months | VT Design Team | (Subject to change once robot and assistive wheelchair have been acquired) | Analysis / PM/M Plan Development
TR 4.0.0 | The system shall be neutral in color and should blend into the motorized wheelchair structure | VT Design Team | Talk this over with customer once robot has been acquired/developed | Visual Inspection
TR 4.1.0 | The system shall contain no sharp edges and should minimize sharp corners on frame, end effector, and joystick | VT Design Team |  | Visual/Physical Inspection
TR 5.0.0 | The robot shall not cause harm to the operator in any way. | VT Design Team |  | Inspection
TR 5.1.0 | Upon startup, the robot shall have safeguards against unexpected energization | VT Design Team |  | Inspection
TR 5.2.0 | The robot shall have physical or program safeguards to prevent it from reaching undesired locations | VT Design Team |  | Demonstration |  |  |  | C | 4
TR 5.3.0 | The robot shall have safeguards against unexpected de-energization | VT Design Team |  | Inspection
Note on G6: • Verification: This column should be populated with a description of the verification document linked to the functional requirement.

## TR 1.1.0 Test Plan
TR # | Requirement Description | Verification Method
TR 1.1.0 | The robot shall be controlled manually by a joystick and should have the ability to program automatic functions accessed through a button push | Test
Test Plan
Step | Instructions | Pass or Fail (P/F)
Overall | Operate all degrees of freedom of the joystick including forward and backward, left and right, twisting the potentiometer, operating the button. Verfiy that the correct movements on the joystick correspond to the correct movement of the robotic arm
1 | Power on the robotic arm
1.1 | Plug in the power cable or batteries to the designated port on the ESP 32
1.2 | Flip the switch to on on the ESP 32
1.3 | Plug in the USB to computer if using python
1.4 | If not using python, ensure the arduino code is uploaded
2 | Verify the +/- x direction movement
2.1 | Push the joystick forward for 5 seconds
2.1.1 | Verify that the robotic arm moves forward from the initial position to its maximum range
2.2 | Move the robot to the initial position
2.3 | Pull the joystick backward for 5 seconds
2.3.1 | Verify that the robotic arm moves backward from the initial position to its maximum range
2.4 | Move the robot to the initial position
3 | Verify the +/- y direction movement
3.1 | Push the joystick right for 5 seconds
3.1.1 | Verify that the robotic arm moves right from the initial position to its maximum range
3.2 | Move the robot back to the initial position
3.3 | Pull the joystick left for 5 seconds
3.3.1 | Verify that the robotic arm moves left from the initial position to its maximum range
3.4 | Move the robot to the initial position
4 | Verify the +/- z direction movement
4.1 | Twist the joystick right for 5 seconds
4.1.1 | Verify that the robotic arm moves up from the initial position to its maximum range
4.2 | Move the robot back to the initial position
4.3 | Twist the joystick left for 5 seconds
4.3.1 | Verify that the robotic arm moves down from the initial position to its maximum range
4.4 | Move the robot to the initial position
5 | Verify the clamping of the end effector
5.1 | The end effector is initially set at 3.14 - closed position
5.2 | Push the button on top of the joystick for 5 seconds
5.2.1 | Ensure that the end effector is fully open
5.3 | Release the button on the joystick
5.3.1 | Ensure that the end effector is fully closed
6 | Power off the robotic arm
6.1 | Remove the USB from computer if necessary
6.2 | Flip the switch to off on the ESP 32
6.3 | Remove the power cord or batteries from the designated port on the ESP 32

## TR 1.3.0 Test Plan
TR # | Requirement Description | Verification Method
TR 1.3.0 | The robot shall have an operating radius that spans a minimum of 5 inches to a maximum of 2 feet from mount location with 180 degrees of rotation total | System Test
Test Plan
Step | Instructions | Pass or Fail (P/F)
Overall | Operate all degrees of freedom of the joystick including forward and backward, left and right, twisting the potentiometer, operating the button. Verfiy that the correct movements on the joystick correspond to the correct movement of the robotic arm
1 | Power on the robotic arm
1.1 | Plug in the power cable or batteries to the designated port on the ESP 32
1.2 | Flip the switch in the ESP 32
1.3 | Plug in the USB to computer if using python
1.4 | If not using python, ensure the arduino code is uploaded
2 | Verify the minimum operating radius of the arm
2.1 | From the initial position, move the end effector as close to [x y z t d] = [0 0 0 3.14 d] as possible
2.1.1 | Ensure that the robot operates normally wihtout any sporatic movements during 2.1
2.2 | Move the end effector to [x y z t d] = [20 0 0 3.14 d]
2.3 | Move the end effector to [x y z t d] = [20 20 0 3.14 d]
2.4 | Move the end effector to [x y z t d] = [0 20 0 3.14 d]
2.5 | Move the end effector to [x y z t d] = [-20 20 0 3.14 d]
2.6 | Move the end effector to [x y z t d] = [-20 0 0 3.14 d]
2.7 | Move the end effector to [x y z t d] = [-20 -20 0 3.14 d]
2.8 | Ensure that the robot revolves around its base in the +z direction to reach the -y plane
2.9 | Move the end effector to [x y z t d] = [0 -20 0 3.14 d]
2.1 | Move the end effector to [x y z t d] = [0 0 0 3.14 d]
3 | Verify the maximum operating radius of the arm
3.1 | From the initial position, move the end effector as close to [x y z t d] = [450 0 0 3.14 d] as possible
3.1.1 | Ensure that the value d does not exceed 449.999999
3.2 | From the current position, move the end effector to [x y z t d] = [0 -450 0 - 3.14 d]
3.2.1 | Ensure that the value d does not exceed 449.999999
3.3 | From the current position, move the end effector to [x y z t d] = [-450 0 0 - 3.14 d]
3.3.1 | Ensure that the value d does not exceed 449.999999
3.4 | From the current position, move the end effector to [x y z t d] = [0 450 0 - 3.14 d]
3.4.1 | Ensure that the value d does not exceed 449.999999
3.5 | From the current position, move the end effector to [x y z t d] = [0 0 0 3.14 d]
4 | Power off the robotic arm
4.1 | Remove the USB from computer if necessary
4.2 | Flip the switch to off on the ESP 32
4.3 | Remove the power cord or batteries from the designated port on the ESP 32

## Definitions
Verification Methods | Example Requirement | Example Verification Process
Inspection | The system shall be green | The customer looks at the system and determines it is green
Demonstration | The system shall not fail immediately when unplugged | The customer unplugs the system when operating and observes that it does not fail
Analysis | The system structure shall withstand a 10g shock load | The team performs an FE analysis with a 10g shock load and presents computational results to support this requirement
Test | The system shall continue to operate for no less than 10 minutes when unplugged | The customer unplugs the system and uses a stopwatch to validate that the system continues to operate for 10 minutes or more
Info from Lecture 2
Verbiage | Approximate legal meaning
Shall | contractually binding, i.e. non-negotiable - if the hardware does not conform, it is not acceptable for delivery to the customer
Should | these are desire-ments, not hard requirements - not meeting requirements in this category would not consititue a non-conformance and the customer will still accept the product. Should be used for requirements that are seen as negotiable as development progresses
Must | A statement of fact to be verified
NOTE: There are a number of other vairations on requirements verbiage. For this class, we will keep it simple and use shall and should to be indicative of threshold and objective targets for quantitative metrics. Binary metrics may use must or shall (e.g. A Class 3B laser safety sign must be affixed to the system)